Summary of Maya Med & Uni Medical
MY-O001B Training

MY-O001B Hemodialysis Machine  —  Installation · Commissioning · Operation · Daily Maintenance

| Field | Content |
| --- | --- |
| Product | MY-O001B Hemodialysis Machine |
| Prepared by | Maya Medical — After-sales & Technical Team |
| Prepared for | Uni Medical |
| Training date | 15 July 2026 |
| Source | On-site MY-O001B training materials: Installation Guide · Operation Training · Daily Maintenance Guide |
| Scope | Installation · Commissioning · Operation · Daily Maintenance |

This document is a condensed reference distilled from the MY-O001B training materials for quick lookup. All values, parameters and procedures are subject to the MY-O001B Product Manual and your local haemodialysis Standard Operating Procedures (SOP). All clinical operations must be performed by qualified medical / technical personnel.

Machine operation is organised into six blocks: Pre-startup preparation → Machine initialization → Patient connection → Monitoring during dialysis → Post-dialysis procedure → Daily maintenance; plus disinfection & infection control, emergency handling, and consumables / water-quality requirements.

# 1. Installation

Installation flow (6 steps):  Unloading → Confirm installation environment → Equipment installation → Power-on test → Functional check → Cleaning.

## 1.1  Unloading & packing list

Open the box, remove the foam, and push the device down the ramp to the floor (do not lift or drag it directly).

Verify the packing list on receipt and record any shortage before signing:

| # | Item | Qty |
| --- | --- | --- |
| 1 | Hemodialysis machine (MY-O001B) | 1 |
| 2 | IV pole | 1 |
| 3 | Fluid supply & return tube | 1 |
| 4 | Equipment manual | 1 |
| 5 | Certificate of conformity | 1 |
| 6 | Warranty card | 1 |

## 1.2  Confirm installation environment (all requirements must be met before installation)

| Item | Requirement |
| --- | --- |
| Power | 230 × (1 ± 10%) V~, 50/60 × (1 ± 2%) Hz; fuse F 15A AH 250V; grounded outlet required |
| Ambient temperature | Operation 10–35 °C; transport/storage −20 to 55 °C (drain & dry the internal circuit if stored below +5 °C) |
| Relative humidity | ≤ 85% (no condensation) |
| Atmospheric pressure | Operation 60–106 kPa; transport/storage 50–106 kPa; altitude ≤ 4000 m |
| Water supply | Inlet 5–35 °C and 5 °C below the dialysate preset temperature; inlet pressure 0.05–0.65 MPa (reduce dialysate flow if at the lower limit); flow ≥ 1.2 L/min; water quality per EN ISO 23500-3:2019 |
| Drainage | 2 drains + 1 floor drain; drain rate ≥ 1.2 L/min; drain outlet ≤ 0.8 m above floor; max drain temperature 95 °C |
| Space | Dialysis unit ≥ 2 × 2 m; power socket every 2 m (≥ 1 m high); drain every 2 m (≤ 0.8 m high); ≥ 10 cm between rear of machine and wall for ventilation |

## 1.3  Equipment installation (parts)

IV pole: place bracket on base → tighten with PH2 screwdriver → loosen holder counter-clockwise → insert hanger, then lock clockwise.

Dialyzer clamp: place onto the IV pole and tighten the black set screw.

Blood pressure monitor: place the cuff on its stand; connect the tube to the BP port on the left side of the machine.

Water pipes: connect the mesh inlet pipe to the water inlet at the rear bottom (RO water inlet) and fasten it to the tap with a pipe clamp (No.5 flat screwdriver); connect the drain pipe to the rear-bottom drain outlet and route the other end into the floor drain.

## 1.4  Before first use

After installation and before first use, perform a citric-acid hot rinse.

# 2. Commissioning

## 2.1  Power on

Open the water valve at the supply port.

Connect the power cord.

Press and hold the power button for 3 seconds to start.

## 2.2  Enter settings mode

Open settings → select [Medical Staff Mode] or [Engineer Mode].

Medical Staff Mode password: 888888.

## 2.3  Functional checks at installation (verify each item)

| Check | Key action |
| --- | --- |
| Peristaltic (blood) pump | Enter [Peristaltic pump]; set the correct tubing diameter |
| Heparin pump | Enter [Heparin]; set parameters |
| Dialysate flow | Enter [parameter]; confirm dialysate flow |
| Dialysate ratio | Enter [Dialysate formula]; select the correct dialysate ratio |

Note: deeper engineer-level calibration (three-pressure zero, conductivity / temperature accuracy, AB pump suction volume, UF pump flow, etc.) is covered in Section 4 · Preventive Maintenance.

# 3. Operation (Use)

## 3.1  Key components

Physical keys below screen: ① Reset  ② Venous-chamber level Up  ③ Venous-chamber level Down  ④ Power  ⑤ Blood pump  ⑥ Mute.

15 panel items: heparin pump; venous / arterial / PBE pressure ports; venous & arterial chamber holders; blood pump; BP monitor; air-bubble + blood detection + venous clamp; arterial monitoring (blood temp / volume); dialyzer connectors (Blue = supply, Red = return); substitution-fluid module; leak sensor; AB suction tube.

## 3.2  Treatment flow (11-step sequence)

Self-test passed → Install bloodlines → Priming → Parameter settings → Ultrafiltration cycle → Confirm parameters → Connect patient → Treatment → Blood return & disconnect → Drain waste → Disinfection.

| Step | Key points & default values |
| --- | --- |
| Power-on self-test | Turn on the rear main switch → tap "Power" → confirm RO water and power are connected. Select "HD Treatment Mode" (auto-defaults to HD after 10 s). Self-test covers circuit, memory, pump, sensors and safety monitoring. Wait for 100% — do NOT operate during self-test; continue only when all items show "Pass". |
| Fluid supply mode | Mode 1: A & B concentrate suction tubes into A & B containers. Mode 2: A concentrate + B dry-powder cartridge. Mode 3: central supply. |
| Change treatment mode | Top-left "Treatment Mode" button; supports HD / HDF pre-dilution / post-dilution / mixed dilution; then "Save". |
| Install bloodlines | After self-test 100% and "Priming" lights up, install the bloodlines and bypass, then tap "Priming". |
| Priming | Tap the blood pump; default 100 mL/min. Priming volume is adjustable; default ≈ 10 min (faster pump = shorter). After priming, an automatic blood-circuit leak check runs ≈ 1 min. |
| Set prescription | "Prescription" → enter treatment parameters; for HDF, set the mode and the substitution target volume. |
| Ultrafiltration cycle | After priming, tap "Recirculation": auto-flush at 200 mL/min ≈ 5 min. Keep the dialyzer upright (red end up) to purge gas; connect the patient only after completion. Note: auto-priming performs pre-ultrafiltration — not applicable when priming with a saline bag. |
| Connect patient | "Connect Patient" → re-verify the prescription → "Confirm". Connect the patient side and tap the blood pump to withdraw blood; default 100 mL/min and 200 mL; auto-stops at target volume or on blood detection. |
| Treatment | "Finished" → treatment screen. After verification, tap the blood pump and set pump speed per the order; the UF pump and substitution pump start automatically. Watch that pressures stay stable. |
| Blood return & disconnect | "Blood Return" → blood pump; default 100 mL/min and 200 mL, linked to the blood detector; auto-stops at target volume or on saline detection. |
| Drain waste | Disconnect the patient, connect the arterial/venous tubing, disconnect the pressure ports → "Empty Blood Circuit" (dialyzer blue end up) → reconnect the blue/red quick connectors to the machine → "Empty Bicartridge", then disassemble. |

## 3.3  Safety principle — Three Checks & Seven Verifications

Three Checks: before, during and after treatment (patient status + operation content).

Seven Verifications: bed number, patient name, dialyzer, dialysate, dialysis time, ultrafiltration volume, anticoagulation method.

Parameters (mode, blood flow, UF volume) must be double-checked against the medical order by two clinicians before being set.

## 3.4  Alarm handling during dialysis

| Alarm | Common causes | Handling |
| --- | --- | --- |
| Arterial pressure | Insufficient blood flow; pump speed exceeds blood flow; air or kink in tubing | Pre-pump infusion; check tubing |
| Venous pressure | Venous access obstruction; venous-chamber filter clog; low flow; pressure fault | Check access / filter. ⚠ Immediate intervention if venous pressure stays > 300 mmHg |
| Transmembrane pressure (TMP) | Dialysate-circuit obstruction; excessive ultrafiltration; dialyzer clotting; venous pressure-shield blockage | Check circuit / clotting |
| Air | Air in lower venous-chamber tube; improper installation; bloodline not connected or ruptured | Check installation & tubing integrity |
| Blood leak | Blood detected; high UF rate → high degassing; drain line loose; oil / calcium deposits | Check drain line & sensor fouling |
| Conductivity (water circuit) | Wrong dialysate concentration; kinked concentrate suction tube; empty concentrate tank | Replace fluid / straighten tube / refill |

## 3.5  Clinical FAQ

Disinfectant dosing (for water-circuit volume 2300): 50% citric acid (2% dilution) = 92 mL; 20% citric acid = 230 mL; 5% sodium hypochlorite (0.1% dilution) = 46 mL.

Dialysate ratio 1 : 1.225 : 32.775; HD dialysate flow 500 mL/min, HDF 700 mL/min.

Does auto-priming/recirculation consume extra A/B solution? No — dialysate flow is constant whether priming online or manually, so total A/B consumption is unchanged.

Pre-/post-dilution: machine default caps pre-dilution blood-flow ratio at 60% and post-dilution at 30% (exceeding triggers an alarm). Per SOP: pre-dilution = 50–60% of blood flow (at 250 mL/min ≈ 30–36 L); post-dilution = 25–30% (≈ 15–18 L).

## 3.6  Emergency handling

Power outage during treatment

Auto-switches to battery with an audible/visual "Mains Power Abnormal" alarm; the water circuit stops, while the blood pump, heparin pump and air detection keep running. A new, fully charged battery sustains these functions for ≥ 30 min (aging shortens or disables this — the battery is a consumable and must be replaced regularly).

On restoration the device returns to its pre-outage state; confirm the extracorporeal circulation is normal before continuing. If power is expected within 20 min, tap "Confirm" after boot to resume.

If the battery is low or power cannot be restored soon: return blood immediately, then power off; after restoration, disinfect promptly to prevent tubing blockage.

⚠ Manual blood return (battery depleted): clamp & disconnect the arterial end → remove the venous end (without the air detector) → connect the arterial line to the saline/substitution port, open the clamp and return blood slowly → open the pump cover and turn the rotor by hand in the arrow direction (clockwise) → after visually confirming all blood has returned, clamp & disconnect the venous end. Always turn clockwise (reverse causes backflow — life-threatening) and turn slowly to prevent air embolism.

Water outage during treatment

Audible/visual "No water or low water pressure" alarm; ultrafiltration stops, the water circuit stops and bypass activates. The Emergency button gives 3-level protection (reduce pump speed → stop ultrafiltration → disconnect bypass); the blood pump auto-drops to 100 mL/min. If water returns within 20 min, use ISO (isolated ultrafiltration) mode to continue while watching pressures; if not, return blood and disinfect promptly after restoration.

Display / touchscreen failure (monitoring & indicators still work)

Press the physical blood-pump key (bottom-right of the screen) to stop the pump → hold the power button 10 s to force shutdown → press 2 s to restart, then choose whether to resume. If unresolved, terminate treatment and return blood via the physical blood-pump key (watch the tubing & patient for air throughout).

## 3.7  Consumables & water quality

Consumables: before use, check package integrity / contamination / expiry; verify model & spec match the order; use aseptic technique (sterile gloves, mask, cap); prevent secondary contamination.

Dialysis water: bacteria ≤ 100 CFU/mL; endotoxin ≤ 0.25 EU/mL; chemical contaminants (chlorine, heavy metals, calcium, magnesium) within limits. Compliant water quality is a precondition for operation — faults caused by non-compliant water are not covered by warranty. Comply with AAMI / ISO 23500 and monitor regularly.

# 4. Daily Maintenance

## 4.1  Cleaning & disinfection frequency

| Frequency | Action |
| --- | --- |
| After each treatment | Thermal disinfection recommended; wipe the housing with 500 mg/L chlorine disinfectant |
| Weekly | At least one custom disinfection: hot citric acid + sodium hypochlorite |
| Weekly | Clean the dialysate connectors / AB suction tubes to prevent crystallization fouling |
| Monthly | Clean the cooling-fan filter |
| Downtime > 48 h | Hot citric-acid disinfection before treatment |
| Service interval | Every 1500 operating hours / 6 months |

## 4.2  Disinfection procedure

Remove the tubing → reconnect the blue quick connector (supply) and red quick connector (return) to the machine → insert the red A connector and blue B connector into their sockets and rotate to lock; ensure the disinfectant suction tube is in the tank with adequate disinfectant → tap "Cleaning" → select mode → default "Thermal Disinfection" ≈ 38 min → choose power-off or self-test after disinfection.

## 4.3  Standby maintenance

Short-term idle: disinfect every 48 hours.

Long-term idle: drain the water circuit; a sample test must pass before reuse; power on & charge the backup battery at least every 3 months.

## 4.4  Consumable replacement cycle

| Cycle | Main parts to check / replace |
| --- | --- |
| 6 months / 1500 h | Fan filter; blue/red supply-return quick-connector seals; A/B suction-connector seals; substitution-fluid assembly seals; dry-powder cartridge seals; button battery |
| 1 year / 3000 h | All of the above + balance-chamber filter, disinfectant suction line, water-inlet filter, A/B concentrate filters, disinfectant filter, backup battery |
| 2 years / 6000 h | Full re-check / replacement of the 1-year set |

## 4.5  Preventive maintenance (engineer)

Precondition: run the system function check first; perform PM only if the system is normal.

System status check (14 items): flushing; citric-acid thermal disinfection (suction volume / heating temperature); self-test; blood-circuit under power interruption; pre-flush; blood / function / substitution pump flow (all at 200 mL/min); degassing; AB & return fluid; blood-leak protection; treatment function; emptying function.

| Module | Key checks / criteria |
| --- | --- |
| Display | Clean; check cables; touch validity; touchscreen calibration (accuracy) |
| Board | Clean; check connectors; three-pressure zero calibration; level & bubble detection; heparin-pump fast-push / diameter / empty-fluid alarm; clamp key; battery charge/discharge; audible & visual alarms |
| Ceramic pump | Clean; check leaks & jams; fixed-speed AB-pump suction volume & calibration; UF-pump flow test & calibration |
| Gear pumps | Clean; front/rear drive-board output 24 V ± 0.2 V; circulation-pump inlet −0.8 kg; front-pump outlet 1.3 kg; rear-pump outlet 1.2 kg |
| Balance chamber | Clean; check terminals/sockets; solenoid-valve open/close; check membrane for rupture |
| Conductivity probe | Clean; check wiring; temperature accuracy; conductivity accuracy |
| Heating chamber | Clean; heating-rod wiring & grounding; heating-rod resistance ≈ 28 Ω; heating function; emptying function |
| Pipeline | Replace AB / inlet / disinfection / balance-chamber filters; water-circuit tightness |
| Venous clamp | Check wiring; switch function; proper clamping |

## 4.6  Maintenance tool list (key items)

Hex-key set; adjustable wrench; hex socket screwdriver (7 mm); Phillips / slotted screwdrivers; diagonal & crimp pliers; vise; shaft/hole circlip pliers; digital multimeter; conductivity meter; rotor calibration tool; soldering iron + lead-free solder; tape measure.

# Appendix — Key Parameters Quick Reference

| Item | Value |
| --- | --- |
| Power | 230 × (1 ± 10%) V, 50/60 Hz; fuse F 15A AH 250V |
| Water inlet | 5–35 °C (5 °C below dialysate temp), 0.05–0.65 MPa, ≥ 1.2 L/min |
| Medical Staff Mode password | 888888 |
| Priming | 100 mL/min, ≈ 10 min |
| Ultrafiltration cycle | 200 mL/min, ≈ 5 min (dialyzer red end up) |
| Blood withdraw / return | 100 mL/min, 200 mL |
| Dialysate flow | HD 500 mL/min; HDF 700 mL/min |
| Dialysate ratio | 1 : 1.225 : 32.775 |
| Pre-/post-dilution cap | Pre 60% / Post 30% (of blood flow) |
| Venous-pressure action line | Immediate intervention if > 300 mmHg |
| Thermal disinfection time | ≈ 38 min (default) |
| Disinfectant (water circuit 2300) | 50% citric acid 92 mL / 20% citric acid 230 mL / 5% NaOCl 46 mL |
| Housing wipe | 500 mg/L chlorine disinfectant |
| Battery backup | Full charge ≥ 30 min (blood pump / heparin pump / air detection) |
| Water quality | Bacteria ≤ 100 CFU/mL; endotoxin ≤ 0.25 EU/mL |
| Service interval | 1500 h / 6 months; downtime > 48 h → hot citric-acid disinfection first |
| Heating-rod resistance | ≈ 28 Ω |
| Gear pumps | Drive board 24 V ± 0.2 V; circulation inlet −0.8 kg; front outlet 1.3 kg; rear outlet 1.2 kg |

This summary was distilled from the MY-O001B training materials for the Maya Medical / Uni Medical after-sales knowledge base. All operations are subject to the MY-O001B Product Manual and local haemodialysis SOP.